## Slide 1: Development of a liquid 3He target system for experimental studies of antikaon-nucleon interaction at J-PARC
Masaharu Sato
RIKEN
The 27th International Conference of the International Nuclear Target Development Society
M. Iio(KEK), S. Ishimoto(KEK), S. Shoji (KEK), S. Enomoto(KEK),
T. Hashimoto(RIKEN), M.Iwasaki (RIKEN) and R.S. Hayano(Univ. of Tokyo)
in collaboration with

## Slide 2: Experiments @ J-PARC K1.8BR
formation (missing mass) &
decay(invariant mass)
Search for antikaon nuclear bound states in the 3He (K-, n) reaction (J-PARC E15)
X-ray detector
X-ray
e-
strong-interaction induced shift & width of atomic level
 Derive information on K-nucleus interaction
1 GeV/c
L
p
p
p-
one of expected decay modes
K-pp  L p p- p p
Kaonic He atom
Spectroscopy of kaonic He atoms (J-PARC E17)
“1st stage data taking in May 2013”
2
Both experiments need liquid 3He target
K-
He
stopped K- reaction on 3He

## Slide 3: Setup
Setup for 3He(K-, n) around the target (E15 setup)
K- @1 GeV/c
n
p,p etc
cylindrical drift chamber (CDC)
3He target
solenoid magnet
beam tracker
3
access only from the downstream side
cylindrical detector hodoscope (CDH)
inner hodoscope (IH)
1.5 m
3He target is install at the center of CDS
PTEP02B011(2012)
Experimental requirements
large size 3He target
(Φ=6.8 cm, L = 12.5 cm)
minimum the amount of the material around the target cell
~ energy loss, multiple scattering
high 6-keV X-ray transmittance of the target cell (for E17).
stability of the temperature
3He density depends on the
temperature

## Slide 4: 3He cryostat
Cover the target with cylindrical detector system
Target cell
TMP
Evaporator evacuation line
separator
(4He, 4 K)
evaporator
(4He, 1.3 K)
heat exchanger (3He, 1.3 K)
4
L-shape cryostat.
K-
side view
~1.2 m
only the target cell is isolated from the main component
NIMA687(2012)1
cooling principle
the target cell & vacuum chamber

## Slide 5: How to condense 3He
Condense gaseous 3He by heat contact with 4He
5
vapor pressure of 4He
separator
(4He,4.2 K)
evaporator
(4He, 1.4~2.5 K)
heat exchanger
(3He,1.4 K
~2.5 K)
4He in
4He gas out
needle valve (NV)
3He gas in
3He liq.
pump (120 m3/h)
latent heat of vaporization
decompression
fins
400 litter of 3He gas (2 x200-litter tanks)
liquid level sensor
(super conducting)

## Slide 6: 3He vapor pressure
vapor pressure of 3He = pressure inside the tank
2.5 K:
p3He ~ 400 hPa
6
1.4 K:
p3He ~50 hPa
achieved temperature ~ amount of the condensed 3He
only 60 % of 3He gas condense
vapor pressure of 3He (and 4He )
(initial ~ 1000 hPa)
separator
4.2 K
To lower evaporator
temperature:
2.5 K (NV open)
1.4 K (NV close)
2.5 K is insufficient
due to the limitation of the total 3He gas amount
NV
* with NV close, liquid in evaporator decreases
3He
4He
* if evaporator is empty, cooling power is lost

## Slide 7: transfer to the cell
liquid transfer between the heat ex. and the target cell
7
3He in the heat ex. flows into the target cell through lower pipe
target cell
~1.2 m
warmed-up 3He returns to the heat exchanger
effective heat transfer by circulation of the liquid between the heat exchanger and the target cell
4He
flow
flow
3He
heat exchanger

## Slide 8: 3He Gas handling system
dry pump
emergency
recovery
3He tank
4He
balloon (1000 litter)
cooling
8
400-litter 3He gas  tank
leak rate : <10-10 Pa m3/s
cell broken!

## Slide 9: photo @ K1.8 BR beam line
side view @K1.8BR
(uninstalled from CDS)
3He target
CFRP (t=1 mm)
transfer tube
500 l 4He Dewar
9

## Slide 10: Beryllium target cell
K-
3He density ~ 0.081 g/cm3
transmittance ~ 85 % @ 5.9 keV
measured with Mn-Ka (5.9 keV)
low amount of material
>99.0 % Be (T=0.3 mm, PF-60)
Al ring
(T = 1 mm)
AlBeMet cap
AlBeMet (Materion Brash)
lower density (2.0 g/cm3 )than Al
higher Young modulus than Al
easy to handling than Be
good x-ray transmittance around 6 keV
c.f. ~50 % with Mylar or Kapton
Thickness VS Transmittance
assembled with EBW
AlBeMet cap

## Slide 11: CFRP vacuum chamber
Why CFRP (Carbon Fiber Reinforced Plastic)?
CFRP is the best choice.
(Beryllium is extremely expensive.)
11
Strength of cylindrical shell ~ Young modulus
Low density is favored from the experiment
CFRP prepreg sheet T= 0.24 mm
(Sankyo Manufacturing Co Ltd. , 80tonUD)
Al frame
(Out of detector acceptance)
4 layers are laminated with 0 and 90 degrees
T =1 mm
beam window
(t = 0.6 mm)
thin vacuum chamber

## Slide 12: Operational result (1)
3He tank open
condense
..saturate
pressure inside the heat ex.
temperature
NV close
condense
refill evaporator
NV open
12
time history of a cooling
close 3He tank valve
NV close
liquid 3He target becomes ready within 6 hours.
no-difference in temperature

## Slide 13: Operational result (2)
Operation parameter
Stably operated in the beam time
refill 4He inside the evaporator
(~ once a day)
13
time history during the beam time of 3.5 days (May 2013)
cell temperature
evaporator liquid level
vacuum level <10-4 Pa
temperature 1.4 K
heat load (1.4 K region) 0.21 W
4He consumption 30 l/day
cooling power : latent heat evaporation of 4He in the evaporator
 need to refill 4He
heat load is evaluated by the reduction rate of the liquid 4He
density fluctuation ~ 0.0810(2) g/cm3

## Slide 14: Target image
K-
reaction vertex :
beam and secondary particle tracks
p,p etc
empty target data
3He target data
confirmed target cell is surely filled up with liquid.
Target image can be checked with beam data
Z [cm]
X [cm]
Y [cm]
Y [cm]
Y [cm]
Y [cm]
Z [cm]
X [cm]
3He target data
Z
Y
X

## Slide 15: Summary
A liquid 3He target system was newly developed for J-PARC E15 and E17 experiments.
All the experimental requirements were accomplished.
Acceptable operational parameters for a long-term (e.g. ~1 month) experiment.
Successfully operated in the J-PARC E15 experiment in May 2013
1st physics paper was submitted to PLB (arXiv[nucl-ex]1408.5637).

## Slide 16: backup

## Slide 17: J-PARC hadron facility
K1.8BR
T1 target (Au)
Tokai, Ibaraki
protons
30 GeV, 30 Tppp (~24 kW) as of May 2013
layout of hadron hall
17
Japan Proton Accelerator Research Complex(J-PARC )

## Slide 18: Beryllium target cell
good X-yay transmittance @ 6 keV
Be cell
pure Be
(T=0.3 mm)
Al ring (T=1mm)
Indium seal
AlBeMet
AlBeMet (trade name of Materion Brash )
composite of Be & Al (powder metallurgy)
K-
Lower density (2.0 g/cm3) than Al
Higher Young modulus than Al
Easy to manufacture than Be
fabricated with EBW

## Slide 19: 3He - 4He heat exchanger
cool down 3He gas with 4He by heat contact
Fin structure (5mm pitch)
Side view
19

## Slide 20: 3He density

## Slide 21: assembly of MICTRON cell
巻き取り機
両面コロナ加工 Toray 20GP46
attach frame & pipe
T=25 mm
接着剤（STYCAST1266）に浸した状態で手巻き （４～５周）
MICTRON成型後 （5巻き、Ｔ＝250～300mm）

## Slide 22: Target cell R&D
critical pressure : > 2 atm (>> normal operation ~ 1 atm)
Geometry
cylindrical (D=XX), L = 10 cm, V= 500 cc  limited by 3He gas volume
material
liquid helium 3 density is low (0.081 g/cm3 @1.3 K)
candidates
 | Mylar | MICTRON | Be
cost | ○ | ○ | ×
X-ray exp | △ | ？ | ◎
In-flight | ○ | ○ | ◎
Mylar cell (T =0.3 mm)

## Slide 23: Transmission of X-rays
Thickness VS Transmittance
Be
MICTRON
PET
C12H10N2O2
Mn Ka
X-ray spectrum by SDD w/ 55Fe
X-ray transmission test
very low transmittance with MICTRON
55Fe

## Slide 24: 3He cryostat
Cover the target with cylindrical detector system
adopt L-shape cryostat.
 Separate the target cell away (~1m) from the major cooling part
Target cell
TMP
Evaporator evacuation line
3He cooling with decompressed 4He
3He leak-tight
Boiling point of 3He (3.7 K at 1atm)
pumping 4He and cooling by latent heat of 4He
3He is extremely expensive!
( > 0.5 M yen / litter ) x (400 litter)
> 2 oku yen
never loss during operation
Separator
(4He, 4 K)
Evaporator
(4He, 1.3 K)
Heat exchanger (3He, 1.3 K)
reduce amount of material around the target
Density of liquid 3He ~0.08 g/cm3
24

## Slide 25: Liquid 4He target
NIM606(2009)233
KEK-PS E549/E570 4He target
cryostat
(side)

## Slide 26: CFRP vacuum chamber (I)
Why CFRP (Carbon Fiber Reinforced Plastic)?
Buckling theory of cylindrical shell (Sauthwell formula)
Pc : critical pressure
E : Young modulus
n : node
l
P
r
High Young’s modulus is favorable
 low –Z material is better to suppress multiple scattering
broken CFRP
n = 2 n = 4
CFRP is the best choice.
(Beryllium is extremely expensive.)
26

## Slide 27: CFRP vacuum chamber (II)
Photo : installed at K1.8BR
forming cylindrical shape
cross
UD
F-W(filament-winding)
prepreg type (sheet-type CFRP, Sankyo Manufacturing Co Ltd.)
 inadequate for thin shell
 Use this type
Tapered vacuum chamber : not to interfere neutron acceptance
CFRP
UD: align fibers in one direction
attach on to a mold by changing fiber direction (0, 90 degree , 4 layers = 1 mm)
carbon fiber / plastic ratio is not high
27

## Slide 28: Pressure test of MICTRON cell
4He gas
cylinder
LN2
P
cell
Pressurize cell with 4He in LN2
T = 250mm (MICTRON 5 turn)  dP > 6 bar
cf. T = 100 mm (MICTRON 2turn) d P > 3 bar
T = 250mm @ 77 K
Dp = 6 bar
耐圧性能を確認