## Slide 1
Maintenance and improvement of beam-line detectors for the BigRIPS fragment separator
Yuki Sato
RIKEN Nishina Center
Detector team

## Slide 2
My work
Maintenance, research, and improvement of the beam line detectors
Task of Detector team
□ Parallel Plate Avalanche Counter : PPAC
□ Multiple Sampling Ionization Chamber : MUSIC
Search of the new detector
□ Diamond detector （SPDR 基礎特研）

## Slide 3
MUSIC and PPACs
MUSIC: DE detector
PPAC: Position detector
PPAC
DE detector
Position measurement
MUSIC
MUSIC and PPACs
are gas detectors.

## Slide 4
Parallel Plate Avalanche Counter : PPAC

## Slide 5
2D position sensitive detector
Fast Amp.
CFD
Pre Amp.
Delay
line (X)
Delay
line (Y)
Particle
Cathode (X)
Cathode (Y)
Anode
Bias
voltage
X1
4.3 mm
X2
Y2
Y1
●iso-C4H10, C3F8, 10～50 torr
●～100V/mm electric field
●Electron avalanche
Delay-Line read-out
H. Kumagai et. al., NIMA 2001
4.3 mm
＋
Parallel Plate Avalanche Counter : PPAC

## Slide 6
2D position sensitive detector
Fast Amp.
CFD
Pre Amp.
Delay
line (X)
Delay
line (Y)
Particle
Cathode (X)
Cathode (Y)
Anode
Bias
voltage
X1
4.3 mm
X2
Y2
Y1
●iso-C4H10, C3F8, 10～50 torr
●～100V/mm electric field
●Electron avalanche
Delay-Line read-out
H. Kumagai et. al., NIMA 2001
4.3 mm
＋
Parallel Plate Avalanche Counter : PPAC

## Slide 7
Photos of PPAC detector
150 mm×150 mm
240 mm×150 mm
Al case
Entrance window
（ Aluminized mylar: 12 mm ）
Pre. amp.
●150 mm×150 mm
●240 mm×150 mm
Prepared

## Slide 8
PPACs are necessary for all experiments.
□ Keep the operational condition.
□ Quick repair is necessary.
(Electrode-foils are expendable item.)
□ Deal with various troubles.
Maintenance
It is necessary to be able to make
all components of the PPAC detectors.
~30 PPACs in BigRIPS and ZDS

## Slide 9
Delay-line fabrication
2 mmφ bakelite rod
Cu wire is wound.
Wound wires are coated
with epoxy resin.
Cu wire coated by polyurethane
Read-out lines are soldered.

## Slide 10
Mounted on printed board.
Chip capacitors (~40 pF)
LC delay-line
Delay-line fabrication
Z0=50 Ω (Matching with input impedance of the pre-amplifier)
Delay：0.81 ns/mm

## Slide 11
Electrode-foil fabrication
Cathode electrode (240 mm×150 mm Y-axis)
4 mm-thick-mylar is pasted and stretched
on the G10 frame.
Evaporation
mask
+
Vacuum metalizing
Delay line was attached.
G10 frame

## Slide 12
Electrode-foil fabrication
Cathode electrode (240 mm×150 mm Y-axis)
+
Vacuum metalizing
Delay line was attached.
G10 frame
2.4 mm
gap:
0.15 mm
4 mm-thick-mylar is pasted and stretched
on the G10 frame.
Evaporation
mask

## Slide 13
Electrode-foil fabrication
Cathode electrode (240 mm×150 mm Y-axis)
+
Vacuum metalizing
Delay line was attached.
G10 frame
Breaking!!
4 mm-thick-mylar is pasted and stretched
on the G10 frame.
Evaporation
mask

## Slide 14
L=240 mm
●： D=3.3 mm
▲： D=4.3 mm
■： D=6.3 mm
Position resolution for 241Am-a particles
Position resolution s (mm)

## Slide 15
σ~200 mm
Reachable position resolutions does not depend on the electrode distance.
L=240 mm
Position resolution for 241Am-a particles
Position resolution s (mm)
L=240 mm
●： D=3.3 mm
▲： D=4.3 mm
■： D=6.3 mm

## Slide 16
Position resolution as a function of the signal intensity
L=240 mm
●： D=3.3 mm
▲： D=4.3 mm
■： D=6.3 mm
(calculated)
Induced charge
(calculated)
Electron drift time
(measured)
：Number of electron-hole pairs
before avalanche amplification
：Townsend coefficient
（by Magboltz simulation）
The most important factor is the S/N ratio.

## Slide 17
Fragment ions (reaction: 238U + 9Be at 345 MeV/nucleon)
Position resolution for heavy ions
H. Kumagai NIMB 2013
C3F8 :10 torr
Anode bias : 1055 V
Signal intensities are increasing.
Z
Position resolution s (mm)
Position resolutions improved.
(Q∝Z2)
s~350 mm at Z≈50 ions.

## Slide 18
240 mm×150 mm Cu cathode
Damage on cathode electrode due to the discharge
Damaged spot
●Electrical discharge was sometimes occur above several tens of kcps beam rate for Z~50 ions.
●Z~10 ions are measurable up to 1 Mcps.

## Slide 19
Solution 1
Changing the electrode materials.
Au, Al, Cu Ag
●Large thermal conductivity : Al < Au < Cu < Ag
●Small electrical resistivity : Al > Au > Cu > Ag
Ag-cathode
(~50 nm-thickness)
The production (evaporation) of the silver electrode is difficult…
Establishment of a manufacturing process is now in progress.
Many spots were appeared.

## Slide 20
Solution 1
Changing the electrode materials.
Au, Al, Cu Ag
● ~50 kcps of Cs (Z=55) ions for 2.5 days.
● ~70 kcps of Z~60-70 ions for ~5 hours.
No damage
Ag-cathode
(~50 nm-thickness)
In the first half-year of 2014…
First test for Ag-electrode (2012/11/6)
● PPAC operation was confirmed using
~160 kcps of Ni (Z=28) ions. (irradiation within 1 hour)

## Slide 21
H.V. Power supply
Signal
separator
Preamp.
+12 V
20 ms Uni-vaibrator
PPAC
Anti-discharge unit
Anode
Cathode
Gas: C4H10 10 torr
10 MW
10 kW
220 pF
Power MOSFET
Anode sig.
Dischage sig.
Development of anti-discharge unit.
Solution 2

## Slide 22
H.V. Power supply
Signal
separator
Preamp.
+12 V
20 ms Uni-vaibrator
PPAC
Anti-discharge unit
Gas: C4H10 10 torr
10 MW
10 kW
220 pF
Power MOSFET
Anode sig.
Dischage sig.
Reduce the damage on electrode foils.
Anode
Cathode
Large discharge current!!
Development of anti-discharge unit.
Solution 2

## Slide 23
Operation test
Example: Discharge induced using 241Am-a by applying the
too large bias voltage, intentionally.
With anti-discharge unit
Without anti-discharge unit
Anode voltage
Cathode
current
Anode voltage
Cathode current
800
700
600
200
100
160
140
120
100
80
60
40
20
Cathode current (mA)
160
140
120
100
80
60
40
20
0
Cathode current (mA)
500
400
300
Anode voltage (V)
800
700
600
200
100
0
500
400
300
Anode voltage (V)
Discharge
growth
0
0
0
-100
-20
0.2
0.4
0.6
0.8
1.0
1.2
1.4
0
Time (ms)
0
0.2
0.4
0.6
0.8
1.0
1.2
1.4
Time (ms)
The anode voltage was forcibly grounded using the anti-discharge unit.
Discharge current was suppressed.
Preventing damages on electrodes.

## Slide 24
Operation test
Example: Discharge induced using 241Am-a by applying the
too large bias voltage, intentionally.
With anti-discharge unit
Without anti-discharge unit
Anode voltage
Cathode
current
Anode voltage
Cathode current
800
700
600
200
100
160
140
120
100
80
60
40
20
Cathode current (mA)
160
140
120
100
80
60
40
20
0
Cathode current (mA)
500
400
300
Anode voltage (V)
800
700
600
200
100
0
500
400
300
Anode voltage (V)
Discharge
growth
0
0
0
-100
-20
0.2
0.4
0.6
0.8
1.0
1.2
1.4
0
Time (ms)
0
0.2
0.4
0.6
0.8
1.0
1.2
1.4
Time (ms)
The anode voltage was forcibly grounded using the anti-discharge unit.
Discharge current was suppressed.
Preventing damages on electrodes.
150×150 mm2 PPAC with the anti-discharge unit was installed at upper stream of F3.

## Slide 25
Summary of PPAC
●The delay-line and electrode-foil are made by hand work.
All PPAC detectors are repaired by detector team.
●The position resolution improves with the signal intensity.
●The durability against electric discharge was improved by
using Ag electrodes.
●Install of the anti-discharge unit for all PPACs
●Optimization of some parameters (Ag-layer-thickness, gas, etc…)
Future work

## Slide 26
Multiple Sampling Ionization Chamber : MUSIC

## Slide 27
Multiple Sampling Ionization Chamber : MUSIC
Beam
Ch.1
Ch.2
Ch.3
Ch.4
Ch.5
Ch.6
20 mm
480 mm
Anode
Cathode
200 mmΦ
232 mmΦ
13 cathodes and 12 anodes (2-m-thick aluminized mylar foils)
Gas: Ar+CH4 (90%:10%), 760 torr
Anode bias 500~550 V

## Slide 28
Multiple Sampling Ionization Chamber : MUSIC
Beam
Ch.1
Ch.2
Ch.3
Ch.4
Ch.5
Ch.6
20 mm
480 mm
Anode
Cathode
200 mmΦ
232 mmΦ
13 cathodes and 12 anodes (2-m-thick aluminized mylar foils)
Gas: Ar+CH4 (90%:10%), 760 torr
Anode bias 500~550 V
Anode
Cathode
-
Heavy ion
-
-
-
-
-
-
+
+
+
+
+
+
+
electron
positive ion

## Slide 29
Multiple Sampling Ionization Chamber : MUSIC
Beam
Ch.1
Ch.2
Ch.3
Ch.4
Ch.5
Ch.6
480 mm
Anode
Cathode
200 mmΦ
232 mmΦ
13 cathodes and 12 anodes (2-m-thick aluminized mylar foils)
Tow anode electrodes are connected to a charge-sensitive preamplifier.
1 channel : 8 cm 6 channels : 48 cm
20 mm

## Slide 30
Energy spectra for heavy ions produced from in-flight fission of 238U at 345 MeV/nucleon.
Average ADC value
Average ADC value
Average ADC value
Average ADC value
Average ADC value
31
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
Average ADC value
The resolving power increased with gas thickness.

## Slide 31
DE
s
○: Z=38
●: Z=51
Consider only energy straggling, then
s ∝ L1/2 (Bohr expression) and DE ∝ L (Bethe-Bloch formula)
thus
s/ DE ∝ L-1/2 .
Energy resolution s/DE (%)
Energy resolution as a function of the gas thickness

## Slide 32
DE
s
○: Z=38
●: Z=51
Consider only energy straggling, then
s ∝ L1/2 (Bohr expression) and DE ∝ L (Bethe-Bloch formula)
thus
s/ DE ∝ L-1/2 .
Energy resolution s/DE (%)
Energy resolution as a function of the gas thickness
Noise level is negligibly small.

## Slide 33
Energy resolution as a function of the atomic number of heavy ions
●: L1=24 cm (3 Ch.)
○: L2=48 cm (6 Ch.)
Consider only energy straggling, then
s ∝ Z (Bohr expression) and DE ∝ Z2 (Bethe-Bloch formula)
thus
s/ DE ∝ Z-1 .
Results of the fitting of Z-1
Energy resolution s/DE (%)

## Slide 34
Average ADC value
Z resolution (DZ)
31
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
Atomic number Z
ADC value ∝ Z2
Width of peaks (σ)
ΔZ

## Slide 35
Z resolution as a function of the atomic number of heavy ions
Z resolutions : DZ ≈ 0.21
L=48 cm
Performance decline is not observed up to Z=52 below 1 kcps.

## Slide 36
Z resolution as a function of the atomic number of heavy ions
ΔZ for uranium is approximately 2 times of that of less than Z=52.
U (Z=92) ΔZ(σ)≈0.45
2014/10/20
Charge state straggling ?
Under consideration
tentatively

## Slide 37
Z resolution as a function of the atomic number of heavy ions
ΔZ for uranium is approximately 2 times of that of less than Z=52.
U (Z=92) ΔZ(σ)≈0.45
2014/10/20
Charge state straggling ?
Under consideration
tentatively

## Slide 38
Ionization chamber (@ F7) Pre. Amp. OUT
PreAmp : Mesytec MPR-16 PreAmp (10 ms decay)
～500 cps
～170 kcps
Many pile-up events !!

## Slide 39
Z resolutions decreased above several tens of kcps.
Rate dependence of Z resolution

## Slide 40
Z
Dependence of energy deposition on atomic number
Recombination of electrons and positive ions is negligibly small.
Decline of DZ was caused by
the measurement system (pile-up events).
DE ∝ Z2

## Slide 41
Summary of MUSIC
「ΔZ」
●Z resolution did not decrease up to Z=52 below, DZ ≈ 0.21
U (Z=92) ΔZ(σ)≈0.45
「High rate」
●Z resolutions decreased above several tens of kcps.
●Recombination of electrons and positive ions was not observed
up to ~100 kcps.
New measurement system is necessary for high-rate conditions.
Charge state straggling ?

## Slide 42
Diamond detector

## Slide 43
Why diamond ?
●Wide band gap (5.45 eV).
●Very high room temperature resistivity (~1016 W・cm)
⇒ low noise operation, no cooling,
●High charge carrier mobility (e:1800 cm2/Vs, h:1600 cm2/Vs)
Si (e:1500 cm2/Vs, h:600 cm2/Vs)
⇒ fast response of detector signals.
●High energy needed to remove carbon atom from lattice (~40 eV)
⇒ detector are radiation hard.
●Using high quality single crystal
⇒ Superior energy resolution.

## Slide 44
CVD diamond detector
e-
h+
e-
h+
e-
h+
Charged particles
e-
h+
e-
h+
e-
h+
e-
h+
Signal
CVD single crystal (4 mm x 4 mm x 90~200 mm)
Metal electrode (Au)
Metal electrode (Al)
weak point

## Slide 45
5.486
5.443
5.389
5.545
Intrinsic energy resolution
~0.4% (FWHM)
at 5.486 MeV
Energy spectrum of 241Am-a-particles
Good energy resolution !!
CVD diamond detector as a
DE detector
Single crystal by Element six.
(4 mm x 4 mm x ~200 mm)
Bias voltage: 100 V

## Slide 46
CVD diamond detector as a
TOF detector
8.6 MeV/u 7Li beam at E7B beam line.
Single crystal by Hokkaido univ. (4 x 4 mm2 crystal)
Collaboration of Detector team, CNS, and Hokkaido univ.

## Slide 47
CVD diamond detector as a
TOF detector
Front side
Rear side
500 ps
FWHM < 1ns
※ oscilloscope: 1 GHz bandwidth

## Slide 48
CVD diamond detector as a
TOF detector
Front side
Rear side
500 ps
FWHM < 1ns
Pulse width
~15 ns (FWHM)
A.Stolz, Expert meeting 2013 at RIKEN

## Slide 49
CVD diamond detector as a
TOF detector
intrinsic detector resolution
σint = 31 ps (σ divided by )
σint = 57 ps
for 58Ni, 600 MeV/u, Z=28 @ GSI
σint = 39 ps
for 238U, 1 GeV/u, Z=92 @ GSI
σint = 27 ps
4He, 8 MeV/u, Z=4 @ CNS
σint = 15 ps
for 78Kr, 87 MeV/u, Z=36 @ MSU
8.6 MeV/u 7Li
s~45 ps
Other Groups

## Slide 50
Summary of diamond detector
● Single crystal diamond detector has a good energy resolution.
● Intrinsic timing resolution is σ~31 ps.
Reported value: 15~57 ps.
NSCL/MSU reported the best value of σ~15 ps.
● Understand radiation damage.
● Fast pre-amplifiers for diamond detector
● Detector with large active areas.
Future work

## Slide 51
Summary of diamond detector
● Single crystal diamond detector has a good energy resolution.
● Intrinsic timing resolution is σ~31 ps.
Reported value: 15~57 ps.
NSCL/MSU reported the best value of σ~15 ps.
● Detector with large active areas.
Crystal growth technology
is making progress day by day.
Direct wafer method
20 mm×20 mm sc-dia.
AIST says “3, 4 inch size wafer in near future”.
20 mm×20 mm single crystal diamond is already commercially available.
Future work

## Slide 52
Summary
Maintenance, research, and improvement of the beam line detectors
□ PPAC
・improvement of the radiation hardness
□ MUSIC
・New signal read-out system
・Study of DZ for heavy ions Z= 60~92.
□ Diamond detector
・large area
・Fast pre-amplifiers for diamond detector
・Study for radiation hardness
Challenges

## Slide 53
fin.