## Slide 1: Recent Advancements and in-house capabilities in Thermal and Mechanical Characterization of Conventional and Novel Materials
Sujit Bidhar
Target Systems

## Slide 2: Outline
Thermal properties : Differential Scanning Caloriemeter,Dilatometer
New techniques-
  Transient Grating Spectroscopy to graphite swelling
  LDV (Laser Doppler Vibrometer) to electron irradiated graphite
New capabilities with Nano-indenter : stress-strain, high strain rate
DIC : Digital image correlation for localized strain

## Slide 3
Differential Scanning Calorimeter
NETZSCH Pegasus F3
Accurate Cp determinations up to 1500°C
Enthalpy precision of better than 0.5%
Operate according to the heat flux principle
Measured values: sample temperature, temperature difference between sample and reference

## Slide 4
Magnetic property change
Change in crystal structure
Melting
Intend to extract dislocation energy and their densities in irradiated material Defect characterization
Defects caused by stress wave
ζ- R method
Thermal resistance and time constant based correction factor to minimize instrument residuals
Iron
Sharp reliable peak, stable baseline
DSC features

## Slide 5
DSC on ceramic nanofiber(ZrO2)
Nanofiber target
Radiation damage resistant
Zirconia Nanofiber Cp by DSC
Measured porosity dependent nanofiber material within 5% error
  developing heat and mass transport simulation in nanofiber media
Useful tool for HEA and validate MD simulation
Working on improving accuracy
Packing density 13.5%
Under 440GeV pulsed proton beam
15%
5%

## Slide 6
Dilatometer
CTE Aluminum A6061
Linseis L75
In vacuum
In Argon
CTE measurement upto 1600 C
dL-rel curve detecting atomic structure change
TGA (Thermo-Gravimetry Analysis) thermal stability
DTA(Differential Thermal Analysis) Determining phase diagram, transition

## Slide 7
Pulsed laser beams, 500 ps, λ500nm, few mJ
SAWs
Interference generate periodic pattern
Absorption of the light within the intensity grating
spatially periodic temperature profile
rapid thermal expansion creates two monochromatic, counter-propagating SAWs
Surface Acoustic Wave (SAW) velocity,
Slow decay low stiffness
Spot Size: 50~500 µm
Transient Grating Spectroscopy (TGS)

## Slide 8: Sensitivity of TGS
10% swelling
50% swelling
Bidhar et al. NME (24), 2020
Lattice swelling and reduction of SAW speed  He filled frankel defect
Hoffman, Acta Mater (89), 2015
Gaussian beam
sigma 1.1mm
Proton beam 120GeV
Peak Fluence: 8.6x1021/cm2

## Slide 9
Limitation:
L =2~3µm in metal, 5~8 µm in ceramics
Low L low signal
TGS
TEM
Correlating TGS-measurable changes with microstructural evolution, it can used as a non-destructive testing (NDT)tool, non-destructive material health monitoring.
Much faster, cheaper than TEM
Thermal diffusivity Vs defects
On Nb, SE Ferry, MIT 2018

## Slide 10: LDV for in-situ target health monitoring
There is clear change in frequency and magnitude of displacement
NTO2 Target Fin
Transient thermal-structural analysis

## Slide 11: LDV on electron-irradiated graphite
Electron Irradiation at SIRIUS (Ecole Polytechnique)
Energy : 2.5MeV
Fluence: 3.4x1019/cm2 ~0.017dpa
Irradiation area : 19x19 mm
10mm
LDV Measurement by Polytec
Excitation : 100 kHz
Piezoelectric at bottom
There is a difference in amplitude and frequency of LDV velocity due to electron irradiation

## Slide 12
Nano-indenter
1000 mN load
Continuous Stiffness measurement (CSM)  E/H Vs Thickness
High strain rate indentation upto 1e4 /sec
Stress-strain curve Determining yield strength

## Slide 13
S
SL
Slope ratio S* = SL/S
Degree of plasticity
0 : perfectly plastic
1: perfectly elastic
*Contact depth
Surface deflection
Oliver Pharr+
Include only 0.8Pmax
Pile up
Sink in
0
Nano-indentation- Young’s Modulus & Hardness
H= P/Ap
Error can be 15~20%!!

## Slide 14
Nano-indenter- Nano Hardness Vs Vickers
Nano-hardness
Force/ Projected Contact Area
Area Calculated from Load displacement
Unit Gpa
Can be small <100nm
More precise
Vickers
Force/ Surface Contact Area
Area Calculated from Optical measurement
Unit kg/mm2
Min indent diagonal >17µm
Operator bais
Lateral contact surface
projected contact surface

## Slide 15
Strain doesn’t change.
Affected volume scales with contact radius and grows during the test
Pyramidal (Berkowich )
Strain changes but affected volume stays constant
Nano-indenter- stress strain curve
ζmax
Spherical
Flat punch
Mean pressure σmean
Uniaxial stress σ
Scaling factor ς
σ = ς σmean
Sneddon’s model

## Slide 16
Aluminum 6061
Modulus by Berkovich indenter
Flat punch
Curve fit σ=kεn
Extrapolate
σy
Avoid substrate effect:
Punch radius < 0.2*thickness of interest
Full contact onset S=2Era
σm = mean pressure = P/contact area
Nano-indenter- stress strain curve
Over estimate size effect
May be accurate in micron size grain
σ = Kεn

## Slide 17
Nano-indentation- high strain rate
Beam induced strain rates are in order of 1000/sec.
Essential to test mechanical properties at these strain rates and mechanism of failure
Loading+unloading in ~100μsec
Failure mechanism is different in high strain rate events : Twinning Vs Slip

## Slide 18
Digital Image Correlation (DIC) : strain measurement
Tensile tester
Camera
Specimen
Aluminum tensile sample
Strain distribution
Validation
Avg strain 0.43%
Strain near hole 1.48%
Kt= 3.4
Valid only for in-plane strain  thin sheets (plane stress)
Developing for full field 3D displacement using stereoscope camera set-up

## Slide 19
Summary
Lab is equipped with some sophisticated instruments
Capability to explore thermal and mechanical properties of thin damaged layer
High strain rate dependent mechanical properties , comparable to beam induced strains, possible
Integrating new non-contact techniques (TGS, DIC,LDV) to evaluate radiation damaged properties

## Slide 20
Thanks for your attention!