## Slide 1: SNS Target Instrumentation Based on Fiber-Optic Sensors: Development and Application Experience
Y. Liu, C. Long, R. Sangrey, and D. Winder
Spallation Neutron Source (SNS)
Oak Ridge National Laboratory
Oak Ridge, TN 37831, U.S.A.

## Slide 2: Outline
Background and motivation
Sensing principle and sensor instrumentation
Measurement results
  Dynamic strain profiles
  Application of measurement
  Application experience
Summary

## Slide 3: Why strain measurements?
Limited space
Strong E-M interference
High radiation dose
High bandwidth
Challenges of strain measurement in the spallation target
Neutrons are produced through spallation reactions between high-energy intense (23 kJ / 700 ns / 1GeV) protons pulses and the mercury.
Proton pulses cause stress cycles in the mercury vessel which lead to fatigue cracks and cavity damage to the structure.
Measurement of the strain waveform on the mercury vessel is critical to improve target performance and lifetime.
Protons
Mercury Target
Strain Sensors
a
b
c
d
e
a
b
c
d
e

## Slide 4: Fiber-optic interferometric sensors
compact
sensitive
immune to electro-magnetic pulse
will not interfere with target leak detectors
Sensor head
~ 65 um
125 um
Fabry-Perot cavity
~ 4 mm
Sensor head
207 um
Advantages of fiber-optic sensors
Single-mode fiber
Reflector
~ 4 mm
Capillary tube
f129 um
f207 um
Incoming light
Thanks to Yasuhiro Takeda (KEK/J-PARC)

## Slide 5: Optical processors for fiber-optic sensors
Opt. Express Vol. 27, 7319 (2019);
J. Lightwave Technol. Vol. 39, 2552 (2021).
Low-coherence interferometry
All-fiber based system
Phase-shifted interference
Digital phase recovery
Self-calibration
Fiber transmission monitor
Fiber Circulator
Fiber Sensor
3x3 FC
PD1
PD2
Faraday mirrors
PD3
Delay line
Laser
Phase shifter

## Slide 6: Instrumentation - sensor installation
Target module
Target assembly
Target vessel
Water cooled shroud
Sensors on target

## Slide 7: Instrumentation - signal processing and data acquisition
Equipment rack
Software platform

## Slide 8: Examples of measured strain profiles

## Slide 9: Validation of target modeling
Measurement
Simulation
Courtesy of J. Hao

## Slide 10: Validation of gas injection effect
Without gas injection
With gas injection
47% strain reduction!

## Slide 11: Strain measurement on STS target samples at LANSCE
  Niobium-clad tungsten block
  Tantalum-clad tungsten block
  Unclad tungsten block
  Validation of numerical modeling
  Assessment of the damping of the strain responses between pulses
  Assessment of the effect of cladding on the dynamic strain response

## Slide 12
~ 2×107 Gy/MWHr
~ 6×105 Gy/MWHr
RIA: radiation-induced attenuation in optical fibers
Sensors stop functioning properly when RIA reaches 20 dB
Radiation effects on fiber sensors

## Slide 13: Radiation effects from single-shot proton pulses
g  1/3

## Slide 14: Summary
We have developed a fiber-optic Fabry-Perot sensor and an all-fiber based signal demodulation system for strain measurement in a spallation mercury target.
The strain measurement demonstrated high stability, high dynamic range, high bandwidth, and high radiation tolerance.
The measurement results on the SNS mercury target have been used to validate the target modeling and the efficacy of strain reduction approach.
Radiation induced effects on sensors have been quantitatively investigated.

## Slide 15
2.3×107 Gy/MWHr
6.4×105 Gy/MWHr
Two pairs of sensors were installed next to each other with different conditions
  Sensors I&F: epoxy applied on the sensor head
  Sensors L&E: no epoxy applied on the sensor head
Static gap of all 4 sensors were monitored as a function of the estimated radiation doses
Radiation effects on epoxy