## Slide 1: SAMURAI Si Detector
M. Kurokawaa), H. Babaa), T. Gunjib), H. Hamagakib),
S. Hayashib),T. Motobayashia), H. Murakamia), A. Taketania),
M. Tanakac), Y. Toganoa), K. Yonedaa)
a) RIKEN, Nishina Center
b) CNS, the Univ. of Tokyo
c) KEK

## Slide 2: Requirements for SAMURAI Si detector system
Coulomb breakup exp.
proton
heavy particle
(z < 50)
Effective area ~ 9 cm × 9cm
Position resolution ~ 1 mm
Dynamic range of the electronics ~ 10^4
  coincident measurement from proton (z = 1) to heavy ions (z ~ 50)  2500
  Pileup of beam events
proton rich particle up to 10050Sn
SAMURAI
Definition of dynamic range
upper limit / lower limit (= 4s)
of detection

## Slide 3: Specification of a Si detector
6 in.
--> 325 mm
single-sided strip detector developed for
GLAST satellite is employed
signals from 3 strips are grouped
together on the PCB
readout pitch = 228 mm × 3 = 684 mm
number of signal channels / detector
= 384 / 3 = 128 ch.
total number of signal channels
= 128 × 5 = 640 ch.
Ref.) T. Ohsugi et al., “Design and properties of the GLAST flight
silicon micro-strip sensors”, Nucl. Instr. and Methods, A541
(2005) 29 - 39.

## Slide 4: PCB design and bonding pattern for Si detector
Si
89.5 mm × 89.5mm
C1
C1
C1
C1
C1
C1
C1
R1
R1
3
5
1
6
C1
C1
C1
C1
stip1
Stip2
Stip3
:
:
:
stip384
stip383
Symmetric layout with upper side
Connectors made by KEL
(USL0040L) for ultrafine coaxial cable
0.4 mm pitch
Outer dimensions = 130 mm × 130 mm
Capacitors to suppress cross talk
through rear side < 1 / 10^4 = 0.01%
enlarged view
Will be ready in May

## Slide 5: Specifications of the electronics
Electronics must have a wide dynamic range ~ 10^4
Electronics must be highly integrated circuit or really tiny discrete circuit
Only conceptional idea exits at the moment
Started development of a CSA (Charge Sensitive Preamplifier)
with a dynamic range of 10^4 in ASIC (Application Specific Integrated Circuit )
less than 2 years ago
For the subsequent stage of the CSA, such as S.A., HINP16C will be used.
Ref.) G. L. Engel et al., “A multi-channel integrate circuit for use in low and
intermediate-energy nuclear physics --- HINP16C”, Nucl. Instr. and Methods,
A573 (2007) 418 - 426.

## Slide 6: pictures of HINP16C
Drs. Yoneda , Baba and I visited
U.S.A. for the negotiation last Nov.
One PCB & one M.B.
will be arrived in Feb. in RIKEN
16 ch/chip × 2 = 32 ch
Mother Board with 16 slots
32 ch × 16 = 512 ch
external preamp. Mode
S.A.  Peak Hold  multiplexer
C.F.D.  hit. pat. & multiplicity

## Slide 7: Cooperative framework for the development of ASIC
RIKEN
KEK
CNS, Univ. of Tokyo
Collaborative research agreement
was concluded on Aug. 1st, 2010
development tool
Experiences
stuffs
Need a CSA with a wide dynamic range
for ALICE calorimeter consisting of W + Si
Fund
modification using FIB (focused ion beam)
specialist of analog circuit: Murakami-san

## Slide 8: Example of the modification using FIB
Disconnection of traces
passivating coating
Insulation layer
Metal (Al) layer
enlarged and tilted view
Limited area was irradiated with
Ga beam
RIKENDGCSP

## Slide 9: Architectures of the present ASIC
Cf = 1.8 pF
Rf = 50 MW equivalent: Krummenacher feed back circuit is used ref.) F. Krummenacher, Nucl. Instru. Meth., A305 (1991) 527.
Power consumption 10 mW/ch
Capacitive division
Openloop gain
Saturation suppression

## Slide 10: Architecture I: dual channel system
Input charge is divided in proportion to the capacitance ratio between coupling capacitances Ch and Cl (Ch > Cl).
Ref. ) N. Uematsu and S. Nishimura, “Development of silicon strip detector with wide-dynamic-range readout system from
20 keV to 4 GeV”, RIKEN Accel. Prog. Rep. 41, (2008) 151.
CSA
CSA
Qin
Ch
Cl
Qh = Qin × Ch / (Ch + Cl)
Ql = Qin × Cl / (Ch + Cl)
Log Qin
Log Vout
4 pC
40 pC
high-gain
channel
low-gain
channel
Extended as much as the
capacitance ratio
Ch = 10 × Cl
H
L
similar response with
single channel system

## Slide 11: Output pulse shapes from the CSAs
Ch
Cl
CSA
CSA
This part is implemented in first prototype
Chip and its test board TSMC 0.5 mm process chip size ~ 1.6 mm × 1.2 mm
H
L

## Slide 12: Result I: Dynamic range for a single channel CSA
ADC-H
ADC-L
10-3
10-2
10-1
100
101
102
100
101
102
103
104
5.2 fc
5.6 pc
Definition of lower limit = 4s. 4.4 fC corresponds to 100 keV.
Dynamic range ~ 10^3
pulser
CSA
SA-HG
SA-LG
ADC-H
ADC-L

## Slide 13: Result II: Dynamic range for dual channel CSAs
10-3
10-2
10-1
100
101
102
100
101
102
103
104
ADC-H-H
ADC-H-L
ADC-L-L
5.8 fc
56 pc
Definition of lower limit = 4s
Dynamic range ~ 10^4
10 times wider than that is obtained
for the single channel usage
56 pF
560 pF
pulser
CSA
CSA
SA-HG
ADC-H-H
SA-LG
SA-LG
ADC-H-L
ADC-L-L

## Slide 14: Summary and time schedule
First prototype of RIKENDGCSP (Riken Dual Gain Charge Sensitive Premplifier) is working well. Dynamic range can be extended as much as the gain ratio.
2009
2010
2011
2012
2013
First prototype of
RIKENDGCSP
second prototype of
RIKENDGCSP
Test with HINP16C
& RIKENDGCSP
Test including Si
detector
Mass production of
circuit and test
first experiment
Good news:
HAMAMATSU has started production of
double sided strip detector.
Selection has increased.