## Slide 1: Summary of INTT ROC Test
2022/11/29
Kazuma Fujiki

## Slide 2: Object
Prepare 16 ROCs with no defects for sPHENIX
16 of 25 ROCs …
▪ are taken over from PHENIX
▪ have malfunctions in many location
▪ need to repair

## Slide 3: ROC Class Definition
Class | Number of the GOOD ports
1 | 12 Ports
2 | 10 or 11 ports
3 | 8 or 9 ports
4 | Less than 7 ports
Not in use for INTT
12 input ports
A1
A2
A3
B1
B2
C1
B3
C2
C3
D1
D2
D3

## Slide 4: Calibration Data From Good Ports
Output ADC
Input pulse height
Input pulse height
Strip channel
Silicon Sensor

## Slide 5: Calibration Data From Bad Ports The problems can be categorized into 4
Calibration Pulse
Symptom : Abnormal calibration data is acquired from all ports or single port
Data Fiber Sync
Symptoms : Some of the LEDs on the FEM do not light up / No data for the corresponding ports.
Power Supply
Symptoms : Abnormal calibration data is acquired / LED behind ROC lights up
Cause : Calibration generation circuit / Multiplexer
Cause : Regulator / Circuit that supplies power to the elements
Cause : Serializer TLK2711 has radiation damage
Data Missing
Symptoms : Some data is missing
Cause : Under investigation
Problem in …
Any unhealthy calibration result indicates the problem of the port.

## Slide 6
Calibration Pulse
Symptom : Abnormal calibration data is acquired from all ports or single port
Cause : Calibration generation circuit / Multiplexer
Multiplexer
Calibration pulse generation circuit
If abnormal data is acquired throughout the column, there is a problem with the calibration generation circuit.
If abnormal data is acquired on a port alone, there is a problem with the multiplexer.

## Slide 7
Calibration Pulse
Symptom : Abnormal calibration data is acquired from all ports or single port
Cause : Calibration generation circuit / Multiplexer
*Trigger signal displayed in cyan
Normal waveform

## Slide 8
Calibration Pulse
Symptom : Abnormal calibration data is acquired from all ports or single port
Cause : Calibration generation circuit / Multiplexer
Trigger Signal
Normal waveform

## Slide 9
Calibration Pulse
Symptom : Abnormal calibration data is acquired from all ports or single port
Cause : Calibration generation circuit / Multiplexer
VOLTAGE FOLLOWER U669 has a low output value
Small amplitude of calibration pulses
400mV
200mV

## Slide 10
Data Fiber Sync
Symptoms : Some of the LEDs on the FEM do not light up / No data for the corresponding ports.
Cause : Serializer TLK2711 has radiation damage
Correspondence between LEDs on FEM and ports
Replacing serializer repaired Fiber Sync in NE5.
The serializer is likely radiation damaged.
25 serializers per ROC
(2 per 12 ports + 1 for FPGA)
Fiber Sync may increase once the experiment starts, especially if the FPGA serializer is damaged, the entire ROC is at risk
Currently approx. 2.7% (8ports/300ports)
2 serializers per port

## Slide 11
  Symptoms : Abnormal calibration data is acquired / LED behind ROC lights up
  Cause : Regulator / Circuit that supplies power to the elements
Power Supply
Calibration data not acquired in column D of NE1
High resistance of column D circuit.
This symptom is explained by the lack of L12, and we have confirmed that it is indeed missing.
  LED behind ROC lights up

## Slide 12
  Symptoms : Abnormal calibration data is acquired / LED behind ROC lights up
  Cause : Regulator / Circuit that supplies power to the elements
Power Supply
Abnormal calibration data acquired at B2 port of NE1
Analog power regulator
Analog power regulator
Digital power regulator
Digital power regulator
At this time, the current value of FPHX is 0.16 A (normally 0.5 A)
B2 (Abnormal)
A2 (Normal)
Output voltage of analog power regulator for B2 is 2.0V lower than normal port

## Slide 13
Data Missing
  Symptoms : Some data is missing
  Cause : Under investigation
Good Data
Whole Chip Data Missing
Likely poor contact between ROC and conversion cable.
Partial Missing Channels
Occurs at side 1 of D3 port with micro coaxial cable.
Worse when wearing the BEX.
Half Entry
Only half the overall number of data comes in on some chips.
Poor contact of pins of DF18 connector / Other
Problem in missing data are categorized into three types

## Slide 14: ROC Test Status at a Glance
15 of 16 Class 1ROCs required
6 ROCs were repaired to Class 1.
One more Class 1 ROC is needed.
2022/11/29

## Slide 15: Status Summary
Class | Quantity
1 | 15
2 | 6
3 | 4
4 | 0
Total 25 ROCs
Code | ROC | Port | Error
C | 1 | 3 | Problem in calibration pulse generating system
F | 6 | 8 | Data fiber doesn’t latch
H | 3 | 5 | Half Entry
P | 0 | 0 | Inadequate power supply to components
D | 3 | 8 | Data collapsed

## Slide 16: Summary
All 25 ROCs were inspected and tested. 34 ports were fixed by repair work.
15 of the 16 Class-1 ROCs needed were ready. One more Class-1 ROC is needed.
Issues associated with partial missing channels or half entry are yet established solution. Keep investigation.
We will continue working on deeper investigation of bad ROC ports and try to fix remaining none class-2 ROCs.
ROC's chart is being prepared with Hikaru.